| ESCAMBIA COUNTY FLORIDA. |
| --- |
| WASTE SERVICES DEPARTMENT |  |  |  |  |  |  |  |  |  |  |  |  |
| SCOPE OF WORK |  |  |  |  |  |  |  |  |  |  |  |  |
| PERDIDO LANDFILL SCALE MAINTENANCE FEBRUARY 2021 THROUGH SEPTEMBER 2021 |  |  |  |  |  |  |  |  |  |  |  |  |

Scope of Services:

The Contractor shall provide all equipment, labor, management, materials, parts, tools and transportation necessary to calibrate, maintain, repair and test the scales at the Perdido Landfill operated by Escambia County Waste Services.

Description of Scales:

Manufacturer:  Fairbanks Scales

- Inbound Scale		Model:  PLT-6010-409

- Outbound Scale	Model:  PLT-6010-408

- Automatic Scale	Model:  PLT-6010-HVB09

Preventative Maintenance

Preventative maintenance will be performed in May and September.  It is scheduled by Waste Services personnel.  The scales at this location contain 24 load cells and 12 controllers.  Maintenance will comply with the requirements of the Florida Weights and Measures Law and the latest edition of the National Institute of Standards and Technology (NIST) Handbook 44.  Preventative maintenance shall include, but is not limited to:

- All traveling expenses to and from the site;

- Visually inspect all components for corrosion and apply proper treatment;

- Check and perform bumper bolt adjustments;

- Inspect and tighten cable connections at load cells and control boxes;

- Inspect and test with certified test weights scales covered in this scope of work;

- Make all normal adjustments to restore and maintain the accuracy of the scales;

- Recommend procedures to identify and prevent defective services from recurring;

- Provide written test results for each scale calibration, as well as, written reporting of any identified issues and recommended repairs:

- Remove cover plates on scales, clean and grease load cells and receiver cups once annually;

- Replace back-up batteries for all inside readers once annually.

Emergency Repair

Emergency repairs will be performed on an as-needed basis.  Emergency repairs include any unforeseen, unanticipated work not listed under Preventative Maintenance.  Initial response to an emergency service request will take place within twenty-four (24) hours.  The response will consist of diagnosing the scale problem and quote for the repair.  Repairs will begin as soon as authorization is given.

Non-Emergency Maintenance and Repair

Non-Emergency maintenance and repairs are those which will be performed on an as-needed basis.  The repairs will take place during normal working hours unless instructed otherwise.  Repairs will begin as soon as authorization is given.  The written quote/response shall include, but is not limited to the following:

- Diagnosis of the problem;

- Cost of parts;

- Travel time for routine repair;

- Estimated hourly labor rate;

- Mobilization of vehicles needed for use of service;

- Warranty information.