{Insert vendor name} PROPOSAL

for Raman microscope (Raman microspectroscopy system) for New Mexico Highlands Univ.

(do not exceed 3 pages)

This procurement is for a Raman microscope (Raman microspectroscopy system) and is funded by a grant from the National Science Foundation, with $150,229 total available funds.

The minimum requirements for this system are:

(i) two lasers (532 nm and 785 nm);

(ii) at least three objective lenses (such as 5x, 20x, and 50x – the latter LWD lens as required by LINKAM for their stage);

(iii) multiple available gratings (variable resolution);

(iv) LINKAM heated/cooled stage (-196 to 600 deg. C range);

(v) PC with monitor;

(vi) Shipping, installation, training etc. must be included in the total price.

NB: Specifications/features below refer to the systems in the proposed configuration (not after-purchase upgrades / options)

Pt 1: Spectral resolution (FWHH; determined using the 15620 cm-1 emission line of neon lamp)

Spectral resolution with the 532 nm laser must be 1 cm-1 or better (0.5 cm-1 preferred). Fill in the table below and briefly comment on the numbers and on the recording conditions used.

| Laser | Spectral resolution (FWHH; determined using the 15620 cm-1 emission line of neon lamp), cm-1 |  |  |
| --- | --- | --- | --- |
|  | XXX gr/mm grating | XXX gr/mm grating | XXX gr/mm grating |
| 532 nm |  |  |  |
| 785 nm |  |  |  |

Pt 2: Spectral range

The spectral range must be at least 50 cm-1 to 3300 cm-1 (ability to go down to 30 cm-1 with the 785 nm laser is highly desirable). Briefly describe how your proposed instrument satisfy this requirement.

Pt 3: System spectral reproducibility (determined based on the variation of band frequency after multiple movements of grating)

System spectral reproducibility of 0.1 cm-1 or better (SD) is needed. State the value and briefly describe how it was obtained.

Pt 4: Additional hardware/software options included and other relevant specifications

Briefly describe / list such options (if any) - for example, enclosure, libraries, etc. Also mention other relevant specifications – for example, spectral stability.

Pt 5: Customer support: application-specific training during installation, service and support afterward

Describe application- specific training (incl. training with the LINKAM stage) of users that will be provided during the installation. Also describe service and support afterwards (including costs associated with such service and support). Please be specific – bullet-point format is preferred.

Append budgetary price quotation (itemized – listing all parts/modules/accessories/S&H/training) and full specifications for the proposed system.